## Slide 1: Removing artifact from the FSCV scan pulses and brain stimulation from the Neuropixels recording
Code is in NPXL_Denoise_ap_bin_file()
In CowenLib

## Slide 2: Why?
Artifact from the scan pulses and especially from the DBS pulses is detected as action potentials and adds hours to the spike-sorting process (and corrupts the process when real spikes get lumped into clusters that map onto artifact).
Thus, as much artifact should be removed prior to spike sorting.
Our tests of common-average and local re-referencing indicates that this does not remove much of the artifact.
The answer: fortunately, we have 384 channels to work with per probe. This gives us an unprecedented capability to ‘model’ the artifact and to identify coincident (presumably artifact) events across channels.

## Slide 3: Similar attempts…
Mendrela, A.E., Cho, J., Fredenburg, J.A., Nagaraj, V., Netoff, T.I., Flynn, M.P., Yoon, E., 2016. A Bidirectional Neural Interface Circuit With Active Stimulation Artifact Cancellation and Cross-Channel Common-Mode Noise Suppression. IEEE Journal of Solid-State Circuits 51, 955–965. https://doi.org/10.1109/JSSC.2015.2506651
Musial, P.G., Baker, S.N., Gerstein, G.L., King, E.A., Keating, J.G., 2002. Signal-to-noise ratio improvement in multiple electrode recording. Journal of Neuroscience Methods 115, 29–43. https://doi.org/10.1016/S0165-0270(01)00516-7
Schelles, M., Wouters, J., Asamoah, B., Mc Laughlin, M., Bertrand, A., 2022. Objective evaluation of stimulation artefact removal techniques in the context of neural spike sorting. J Neural Eng 19. https://doi.org/10.1088/1741-2552/ac4ecf
Young, D., Willett, F., Memberg, W.D., Murphy, B., Walter, B., Sweet, J., Miller, J., Hochberg, L.R., Kirsch, R.F., Ajiboye, A.B., 2018a. Signal processing methods for reducing artifacts in microelectrode brain recordings caused by functional electrical stimulation. J. Neural Eng. 15, 026014. https://doi.org/10.1088/1741-2552/aa9ee8
Young, D., Willett, F., Memberg, W.D., Murphy, B., Walter, B., Sweet, J., Miller, J., Hochberg, L.R., Kirsch, R.F., Ajiboye, A.B., 2018b. Signal processing methods for reducing artifacts in microelectrode brain recordings caused by functional electrical stimulation. J. Neural Eng. 15, 026014. https://doi.org/10.1088/1741-2552/aa9ee8

## Slide 4
The figure on the left indicates data with a local car. The FSCV scan pulse is near the center of the image.
The result of using the regression approach to reduce artifact is indicated on the right – same time window. The artifact is nearly absent. I was quite impressed.

## Slide 5
The situation for the DBS artifact is more serious. The artifact is much larger.
The left artifact is the FSCV scan, the right is DBS. The artifact removal process 1) does the regression approach to model the artifact and remove it (later in the slides). 2) stuff that was not removed using this procedure (defined as an event 3 standard deviations above baseline simultaneously on >20 channels) was then turned to zero.
This does work, but leaves a lot of squarish artifacts relating to the zeroed out periods. An improvement could be to merge artifact windows within say 10 msec of each other. This would not be hard.

## Slide 6: Procedure:
Perform common average re-reference where the median over time is subtracted from each channel and then the median for each time-slice is also subtracted.
Create a training and test set from data sampled throughout the entire recording episode. (a 384 x n samples matrix).
Iteratively go through each channel in a for loop. For the current channel y_i, predict the activity of y_i through regression using adjacent channels as predictor variables (Ivs in the regression). Restrict the adjacent channels to be at least 5 channels away from the target channel y_i so that we don’t subtract away the spikes we want to preserve. Set an outer window as well of < 23 channels so that the regression is local and less complex. Train the model on the train set, test on the test set. Use RMS and mean RMS across channels on the test set as the metric used to assess performance.
Apply this model to the entire dataset. Predict the activity on each channel and replace the channel data with the RESIDUALS of the prediction. That is the new signal.
Detect coincidences in the now cleaned data. Define coincidences as signals > 3 standard deviations above the abs value of the data that occur simultaneously on at least 23 channels.
Model these events using regression as was done above and subtract. Hopefully this reduces some of the noise and fills in some data.
For data that still crosses the 3sd threshold on > 20 channels, turn that data to zeros.
This process is slow and on my rather fast computer using a SSD it takes almost 3x times the duration of the recording to run.

## Slide 7: Lessons learned
Linear regression appears to work the best relative to stepwise regression, SVM, simple quadratic regression, and full quadratic. The latter two have so many parameters that it they over fit the data. I was surprised that linear does better than SVM and simple quadratic.
  It is conceivable that more training data could improve the model is this could reduce the overfitting – maybe.
  A neural net might be good too? Probably not worth the inscrutability of the network parameters though. Linear regression is easy to interpret.
  We could also do lasso regression to find a more compact model (similar to stepwise).
  Stepwise was very slow to develop the model.